## Sheet1
All Reported Data are CDS measurements, found under e: \Data-backup\HXRG\CDSReference for earlier than 10/13/09, and under c: \Data\HXRG\H2RG-ASIC-6\CDSReference later,
unless labeled F-i, in which case they are under FSRamp rather than CDSReference. Data are in 4502 configuration unless noted, and are in 4 output mode.
Data are reported for a gain of 4, converted from whatever gain they were measured to a gain of 4. There are 3.84 e-/ADU at gx4 for the spectrometer array.
Noise @gx4 | [0,36], [37,511] -1 | Output-2 | Output-3 | Output-4 | Vref1, Vrefmain, Vbiasgate | Comments | Single frame
| (ADU) | (ADU) | (ADU) | (ADU) | (V) |  | (ADU)
Data Directory
\20091009132747 | 7.0, 6.1 | 7.0, 6.2 | 7.0, 6.3 | 7.0, 6.4 | 1.25, 0.90, 2.05 | silver box, analog/dig ground connected, Ipreampbias=5.4ua | 34500
\20091009133530 | 6.7, 6.0 | 6.6, 5.9 | 8.0, 6.9 | 6.8, 6.4 | “ | Ipreampbias = 6
\20091009171159 | 7.0, 6.2 | 6.7, 6.0 | 8.1, 7.1 | 7.0, 6.5 | “ | Ipreampbias = 8
\20091009171612 | 6.8, 6.1 | 6.5, 5.8 | 7.9, 6.9 | 6.8, 6.4 | “ | Ipreampbias = 10
\20091009171932 | 11.1, 9.4 | 9.9, 8.5 | 12.7, 11.0 | 11.0, 8.6 | 1.25, 1.2, 2.05 | Ipreampbias = 10 | 12000
\20091009172413 | 12.8, 11.1 | 11.47, 9.9 | 15.2, 13.3 | 12.7, 10.0 | 1.25, 1.3, 2.05 | Ipreampbias = 10 | 5000
\20091013113124 | 6.8, 6.0 | 6.5, 5.9 | 8.0, 7.0 | 6.7, 6.4 | 1.25, 0.9, 2.05 | silver box, analog/dig ground connected, Ipreampbias=5.4ua | 34630
\20091013113702 | 5.9, 5.2 | 5.8, 5.0 | 7.2, 6.1 | 5.8, 5.5 |  | actual gx8 (actual data are divided by by 2) | 37000
\20091013141209
F-1 | 7.8, 6.9 | 7.8, 6.3 | 8.7, 7.6 | 8.9, 7.2 | “ | actual gx4 – Fowler frames
F-2 | 6.1, 5.5 | 5.3, 4.9 | 6.4, 5.7 | 6.1, 5.5
F-4 | 4.9, 4.4 | 4.0, 3.8 | 4.7, 4.2 | 4.4, 4.2
F-8 | 4.1, 3.6 | 3.2, 3.1 | 3.6, 3.2 | 3.6, 3.3
\20091013151342
F-1 | 7.3, 6.2 | 7.3, 6.3 | 7.3, 6.4 | 7.3, 6.5 | “ | actual gX11.3 (actual data are divided by 2.825) | 38500
F-2 | 5.6, 5.0 | 5.2, 4.5 | 5.8, 5.1 | 5.6, 5.1
F-4 | 4.6, 4.1 | 4.1, 3.8 | 4.5, 4.0 | 4.6, 4.1
F-8 | 3.9, 3.5 | 3.2, 3.1 | 3.5, 3.1 | 3.8, 3.4 |  | sqrt N = 2.83
\20091013153958
F-1 | 7.2, 6.3 | 7.2, 6.4 | 7.2, 6.5 | 7.2, 6.6 |  | actual gx16 (actual data are divided by 4)
F-2 | 5.6, 4.8 | 5.5, 4.7 | 6.6, 5.1 | 5.8, 5.0
F-4 | 4.7, 4.1 | 4.1, 3.5 | 4.5, 3.9 | 4.8, 4.1
F-8 | 4, 3.6 | 3.1, 2.9 | 3.5, 3.1 | 4.0. 3.4 |  | nothing like sqrtN reduction and nothing like gain advantage
\20091008122900 | 6.6, 6.1 | 6.5, 5.8 | 7.8, 6.9 | 6.6, 6.4 | 1.25, 0.909, 2.05 | gx4 checking noise as a function of single frame baseline | 34000
\20091008124303 | 7.2, 6.3 | 6.7, 6.0 | 8.2, 7.0 | 7.1, 6.5 | 1.25, 1.0, |  | 27400
\20091008124748 | 8.3, 7.4 | 7.5, 6.7 | 9.2, 8.2 | 8.2, 7.2 | 1.25, 1.1 |  | 19900
\20091008125239 | 10.2, 9.0 | 9.0. 8.0 | 11.4, 10.2 | 9.8, 8.2 | 1.25, 1.2, 2.05 |  | 12200
\20091008130136 | 12.0, 10.7 | 10.6, 9.4 | 13.8, 12.4 | 11.7, 9.6 | 1.25, 1.3, 2.05 |  | 4800
\20091008132616 | 12.1, 10.8 | 10.7, 9.5 | 14.0, 12.5 | 11.6, 9.6 | 0.75, 0.8, 2.05 |  | 5100
\20091008140507 | 8.9, 7.7 | 8.2, 7.4 | 10.8, 9.6 | 8.7, 7.9 | 1.25, 0.7 | there are others did not include all – same trend | 49000
\20091008141505 | 7.3, 6.5 | 7.0, 6.3 | 8.7, 7.8 | 7.2, 7.8 | 1.25, 0.8 |  | 42000
\20091008093945 | 6.8, 6.1 | 6.5, 5.8 | 7.7, 6.7 | 6.8, 6.4 | 1.25, 0.96 |  | 30660
\20091008094931 | 6.8, 6.0 | 6.5, 5.8 | 7.9, 6.8 | 6.8, 6.3 | 1.25, 0.909, 2.05 | effect of Vbiasgate | 34250
\20091008095624 | 8.0, 7.8 | 6.6, 7.7 | 11.8, 10.9 | 7.8, 9.2 | 1.25, 0.909, 2.078 | pic frame bright not dark, contrast down
\20091008100311 | 7.5, 7.4 | 7.2, 6.2 | 10.6, 9.6 | 7.5, 8.2 | 1.25, 0.909, 2.065 | pic frame still bright
\20091008101815 | 7.2, 6.8 | 6.9, 6.0 | 9.5, 8.6 | 7.4, 7.2 | 1.25, 0.909, 2.04 | pic frame normal
\20091008101344 | 7.1, 6.7 | 6.7, 5.9 | 9.3, 8.3 | 7.1, 7.3 | 1.25, 0.909, 2.01
\20091008101830 | 6.7, 6.0 | 6.5, 5.7 | 7.9, 6.9 | 6.6, 6.4 | 1.25, 0.909, 1.95 |  | 32000
\20091008102335 | 6.7, 6.0 | 6.5, 5.9 | 7.1, 6.5 | 6.8, 6.3 | 1.25, 0.909, 1.9295 | larger 37 col percentage wise at lower noise | 33000
evaluating gain below, same params as line above
\20091008103250 | 5.7, 5.1 | 5.9, 5.1 | 7.2, 5.7 | 6.1, 5.4 | 1.25, 0.909, 1.9295 | gx8 (divide by 2, all else same as prior entry) | 33100
\20091008103729 | 5.7, 4.9 | 5.6, 4.8 | 6.9, 5.5 | 5.8, 5.2 | 1.25, 0.909, 1.9295 | gx11.3 (divide by 2.8275) – DID NOT GET GAIN ADVANTAGE – checked lab calc – check files
\20091008104631 | 5.7, 4.8 | 5.6, 4.7 | 7.0, 5.5 | 5.7, 5.1 | 1.25, 0.909, 1.9295 | gx16 (divide by 4) | 34000
grounding tests | all the above are with dewar floating, cable shield (on cable from Jade to dewar) floating, external power supply (analog) negative grounded to enclosure case, USB (computer ground) and
| power supply ground, in common AC receptacle. |  |  | Analog and digital grounds connected at Jade board (w3 jumper in).
\20091013164709 | 11.5, 4.7 | 8.5, 4.8 | 15.2, 4.9 | 10.9, 5.0 | All are gx16 (divide by 4) | cable shield and dewar connected to enclosure case, 5V supply grounded
\20091013171153 | 10.8, 4.7 | 8, 4.8 | 14, 4.9 | 10.3, 5.0 | 1.25, 0.93 | w3 out | 32100
\20091013172149 | 5.6, 4.8 | 5.5, 4.6 | 6.8, 5.6 | 5.5, 5.1 |  | w3 out, dewar floating, cable grounded to enclosure, 5V supply grounded to case
\20091014131859 | 5.9, 5.1 | 5.3, 4.6 | 7.0, 6.0 | 5.4, 5.2 |  | 5V power supply floating, w3 out, dewar and cable shield floating, 5 V supply grounded to case
\20091014132544 | 14.5, 12.5 | 25, 12.2 | 24.5, 13 | 13.2, 15 |  | w3 out, dewar and cable shield grounded to box, 5V supply floating
\20091014135002 | 21, 21 | 18.3, 17.8 | 36.3, 35.8 | 18.3, 18.5 |  | Repeat previous with extra ground connection to water pipes
\20091014161939 | 6.3, 5.5 | 5.8, 5.0 | 7.5, 6.5 | 6, 5.8 |  | w3 open, 5V supply floating, dewar floating, cable connected to dewar
\20091014162511 | 6, 5.5 | 5.5, 5.3 | 7, 6.8 | 5.8, 5.5 |  | Cable shield connected to enclosure otherwise same
\20091014163503 | 5.9, 5.5 | 5.7, 5.3 | 6.9, 6.7 | 5.7, 5.6
\20090924095855 | 6.0, 5.5 | 5.6, 5.4 | 6.8, 6.6 | 6.0, 5.5 | 1.7, 1.5, 2.05 | gx4, InpCom on V1, V2 and Vref1 V3,Vrefmain V4 (usually PadP on V2)=4522 | 48200
“input shorted” - showed previously that the difference in Vref1 and Vrefmain matters, not Vref1
\20090916 | Power supply and Jade board in metal enclosure.
|  | standard | best |  |  | Best 4.6 ADU = 17.7 e-
| UR controller | UR Well size | ASICgx4 (e-) | ASIC Well size
F-1 | 22 e- | 143000** | 23.04 | 122880
F-8 | 9.5 e- | e- | 11.52 | e-
** 132000 corrected for IPC
100,000 e- goal